About This Document

The Forensic Lab Director / Quality Manager reviews this document at least annually.  If changes are made, analysts acknowledge the updated procedures.  Obsolete procedures are archived and retained in the laboratory for at least two years.

TABLE OF CONTENTS

- INTRODUCTION

  - Principle

  - Specimen Requirements

  - Chemicals

  - Safety Precautions and PPE

  - Apparatus

  - Preparation of Solutions

  - QC and Sample Run Scheme

- Evidence / Sampling Handling and Preservation

- Quality ASSURANCE

  - Equipment Maintenance and Calibration

  - Reagents, Standards, and Quality Control Materials

  - Quality Control

  - FREQUENCY of Updating calibration curve

  - evaluation and reporting results outside the range of   calibration

- Sample Preparation

- Gas Chromatographic Analysis

- Calculation and Reporting of Results

- Notes

- Case Documentation

  - Case Notes

  - Case File

- Measurement of Uncertainty

## INTRODUCTION

Forensic alcohol analysis is defined as the practical application of specialized devices, instruments and methods by trained laboratory personnel to measure the concentration of ethyl alcohol in samples of blood from persons involved in traffic accidents or traffic violations.

This activity is carried out in laboratories certified by the Maine Department of Health and Human Services (D.H.H.S.).  In order to be so licensed, personnel must meet the requirement set forth by DHHS rule 10-44 Chapter 267 - CERTIFICATION STANDARDS FOR PERSONS CONDUCTING CHEMICAL ANALYSES OF BLOOD AND BREATH FOR THE PURPOSE OF DETERMINING THE BLOOD ALCOHOL LEVEL.

The Maine Health & Environmental Testing Laboratory meets the above criteria and is a certified forensic alcohol laboratory.  The method selected for the determination of alcohol content of blood samples utilizes a headspace gas chromatograph to perform a test that is both qualitative and quantitative.  The procedure calls for addition of a small aliquot of sample to an internal standard solution.  A portion from the headspace of this mixture is injected onto gas chromatographic columns that are capable of separating ethyl alcohol from acetone and the common aliphatic alcohols (i.e., methanol, isopropanol, etc.).  Quantitation is accomplished through comparison to calibration curves.  Data is captured and calculations are performed by device(s) designed to do so (i.e., integrator, workstation, laboratory automation computer).  The method is available to all via SharePoint, where the lab stores controlled copies of documents, forms, etc.

## PRINCIPLE

Aliquots of biological fluids or liquids are mixed with an internal standard solution.  The samples are then analyzed by headspace gas chromatography and quantitated using the internal standard technique.

## SPECIMEN REQUIREMENTS

    - Whole blood-serum-plasma

    - Urine

    - Liquids and/or beverages (not on the Lab Scope of Analysis)

  - Chemicals

Deionized water-ethanol standards-volatiles standards-whole blood controls

## Safety Precautions and PPE

Lab coats, gloves and eye protection will be worn when handling chemicals.

Full-face shield will be worn when handling blood samples.

  - APPARATUS

Model 7694 Agilent Headspace Sampler

Agilent 7890A Hewlett-Packard gas chromatograph

Detector – Dual Flame Ionization (FID)

Columns – Restek RTx – BAC 1; 30 meter; 0.53mm ID; 3.0 um df

Restek RTx – BAC 2; 30 meter; 0.53mm ID; 2.0 um df

Carrier Gas – Helium UHP

Detector Gas – Hydrogen Zero

Headspace Parameters – 	Shake vial			1min. (slow)

Stabilize time		1 min.

Vial pressure		10.6 psi

Carrier pressure		12.4 psi

Oven temp.		50oC

Loop temp.		80oC
				Transfer line temp.	80oC

Cycle time		4.1 min.

Vial equil.		10 min.

Pressurization time	0.13 min.

Loop fill time		0.15 min.

Loop equil. Time		0.15 min.

Inject time		0.20 min.

Chain methods		1+1

Check GC Ready		No

GC Parameters -			Oven temp.		40oC

7890A				Inlet temp.		180oC

Inlet Pressure		24 psi

Detector temp. A&B	240oC
				Inlet B			180oC

Init. Temp		40oC

Init. Time		3.50 min.

Oven Equil. Time		0.50 min.

Oven Max.		225oC

Ambient			25oC

Oven			ON

Init. Value A&B		OFF

On Time A		0.75 min.

On Time B		0.00 min.

Off Time A		0.00 min.

Off Time B		0.00 min.

Save Data		BOTH

Detector A&B		ON

Peak Width		0.039 Col A / .078 Col B

Peal Integration

Data Rate		5.000

Data Storage		ALL

Attn. Sig. 1		3

Attn. Sig. 2		4

Offset Sig. 1		10

Offset Sig. 2		10

Time Sig. 1		5.0

Time Sig. 2		5.0

Quant Report		Summary/Printer

  - Preparation of Solutions

Internal Standard:  0.02% by weight n-propanol

Dilute 200.0 uL n-propanol to 1.0L with deionized water

Ethanol Calibrators and Standards:

Purchased from a approved vendor (such as Cerilliant / Lipomed, etc)

Whole Blood Controls from an approved vendor (such as CliniQA Laboratories, etc)

Volatiles:  Volatile calibrators shall be purchased from an approved vendor (such as Cerilliant,  Lipomed, etc).  Daily samples containing volatiles are part of the whole blood controls from CliniQA Laboratories.

  - QC and Sample Run Scheme:

Case samples are run in duplicate.  The following is an example of how  casework samples and QC checks/standards may appear on the batch sheet:

Water Blank

ETOH Standard

3 samples in duplicate

ETOH Standard

3 samples in duplicate

ETOH Standard

3 samples in duplicate

ETOH Standard

3 samples in duplicate

ETOH Standard

3 samples in duplicate

ETOH Standard

2 samples in duplicate

ETOH Standard

Whole Blood Control – Low BAC

Whole Blood Control – High BAC with Volatiles

The form used to record such information can be found on SharePoint.

- Evidence Handling and Preservation

All laboratory personnel will handle submitted materials in a manner that assures the integrity of the evidence.  Prior to initiating and during the processing of evidence, the analyst will employ the following practices:

- The work area will be clean and free of any excess debris

- Countertops will have adequate space for working with samples

- All glassware and tools to be used will be clean

- Test tubes, capillary pipettes, Pasteur pipettes, etc are used only once, then discarded

- To prevent cross contamination of samples, only one case will be opened by the analyst at a time

- Reagents and solvents will be kept in closed containers when not being used in the analysis

During analysis the evidence will be under constant control by the analyst.

Evidence to be analyzed will be removed from evidence refrigerator and the reverse side of the pink Receipt/Request for Examination Form will be filled out (i.e., internal chain of custody)

The analyst will initial stickers bearing the Lab Identification Number.

If the subject’s name is not available at the time of log-in, the analyst will write the subjects name on the label at the time of analysis (if known).

The analyst will ensure all identification numbers and names agree with the Chain of Custody receipt.

The collection kit and all specimens will be labeled with the lab identification number, name of the subject (if known) and the analyst’s initials.

All paperwork contained in the kit will be labeled with the laboratory identification number and initialed by the analyst.

The analyst will verify the case information provided with the kit matches the HETL folder, sample information from the Blood Alcohol Analysis form submitted with the sample and all Starlims labels.  Any discrepancy shall be noted by the analyst within the case notes, and/or batch sheet(s).

The analyst will document the HETL case number on the worksheet.

The analyst will record the lot numbers of the standards, control and calibrators on the worksheet

At the time of analysis a worksheet with the specimen identification number will be created.

After analysis the remaining blood tubes will be sealed in an evidence bag, the seal initialed by the analyst, and the bag labeled with the laboratory identification number.  The bag is placed in the appropriately labeled sample bin located in the evidence refrigerator.

The sample kit (everything other than blood tubes)  will be stored in an appropriately labeled box.  This box will be retained until filled.  All filled boxes will be placed in storage until returned to the submitter or destroyed.

- III Quality Assurance

A. Equipment Maintenance and Calibration

Refer to Quality Manual

B.  Reagents, Standards, and Quality Control Materials

Refer to Quality Manual

C. QUALITY CONTROL

Function checks will be performed to check the performance of equipment and agents used (either at regular intervals or while testing samples).

Control checks will be performed during the analysis or testing process. These checks are used to:

Determine the performance of the analytical or testing system.

Quantitate the variability of results from the analysis or test in terms of precision and accuracy.

The frequency of checks will be determined by:

- Currently accepted practices/standards in the discipline.

- The number of samples being run in a particular sequence.

Wherever possible, Control Charts will be set up and used to record results from selected function and control checks.  Determination will be made whether the testing or analytical process is out of control and corrective actions taken will be recorded.

Control checks will be performed during the analytical or testing process.  These checks are performed either with each analysis or intermittently after a specified number of analyses.  These control checks include but are not limited to:

- Blanks

- Standard with known or established specifications

- Running samples in duplicate

Calibrators and standards shall be from different sources.  Lot numbers and expiration dates will be recorded on worksheets.

Ethanol Calibrators and Standards will be traceable to Guide 34 compliant suppliers whenever possible.  The Forensic Lab Director / Quality Manager will maintain appropriate records of approved vendors, and how vendors are approved.

- FREQUENCY OF UPDATING THE CALIBRATION CURVE

A calibration will be updated at least once per week when samples are analyzed.  This data will be stored with the ethanol control documents

Criteria for acceptance will be:

1) A minimum of 4 points; and

2) An   r2 value of .998 or greater

E. 	Evaluation and reporting results outside the range of the calibration curve

Any sample determined to be greater than the largest calibrator will be reported as ‘greater than 0.xxx’  (Where xxx denotes the concentration of the largest / highest calibrator).

- SAMPLE PREPARATION

1. Mix the sample thoroughly

2. Label two 20mL headspace vials with identification number and suffix (A or B).

3. Using the auto dispenser re-pipetter - pipette 2500 uL of  internal standard solution into each vial.

4. Pipette 250ul of sample into each vial.

5. Seal the vial by crimping the vial cap.

6. Vortex the vial.

## GAS CHROMATOGRAPHIC ANALYSIS

- Check helium and hydrogen tanks and replace if necessary.

- Load vials on headspace autosampler.

- Verify the correct order of samples placed in the sample tray;

- Set vial parameters on headspace control panel (first vial, last vial).

- Store method 1.

- Load method 1.

- Check carrier pressure should be ~12.4 psi.

- Turn on hydrogen at the main valve on the tank.  On QC console press “Front Detector”.

Use the arrow buttons to move the cursor down to “Flame”.  Press the “on/yes” button to        light the detector.  Repeat the process for the back detector.

- Check detector signals on GC and re-light if necessary.

- Load ETOH250 method on chemstation on computer.

- Create a run sequence by entering the sample numbers in the sample log table.

- Save the sequence in the form of C:\MSDCHEM\3\SEQUENCE\mmddyy.

- Save the data path in the form of C:\MSDCHEM\3\DATA\mmddyy.

- Simulate the sequence and say yes to view it and print it.

- Press start button on headspace control panel

- Run the sequence.

- At the conclusion of the run shut off the detector flames and shut the valve on the

hydrogen tank.

NOTE: Make notation on the worksheet of any instrument repair and/or removal of any vials (specify sample numbers) which occur during the analytical run, or any issue that prevents the run from going to completion.

- CALCULATION and REPORTING OF RESULTS

The acceptable limits of accuracy for the standards, sample replicates during the run are as follows: + 0.005 g/dL or + 5%, whichever is greater. (Samples less than 0.010 g/dL will be reported as ‘less than 0.010 g/dL’ even if difference in duplicates is greater than .005 g/dL) To better define and clarify exactly what is acceptable and what is not: There are 4 results from each case (duplicate samples examined by dual column FID). Samples will be compared as such: Vial A: Result 1 to Result 2: Vial B: Result 1 to Result 2: Vial A: Result 1 to Vial B: Result 1: Vial A: Result 2 to Vial B: Result 2: CLARIFICATION FOR STANDARDS: There are 2 results from each QC Check-Standard. (Result 1 and Result 2). Results will be compared as follows: Result 1 to Result 2 Result 1 to nominal value of Standard Result 2 to nominal value of Standard For both Casework and Standards: Acceptable results are less than or equal to 5.0%, or 0.005, whichever is greater. Meaning, 5.01% = Fail. 0.005 = PASS If any sample from casework fails, at least 1 (2 if adequate sample exists) new vial(s) will be prepared and tested along with at least 1 blank, 2 aqueous controls, and 1 whole blood control. If any QC Check-Standard fails, then all casework samples that bracket the failure will be re-aliquoted and tested, with at least 1 blank, 2 aqueous standards, and 1 whole blood control. WHOLE BLOOD CONTROLS: Whole blood controls (low and high) are examined in duplicate and compared to each other and to the range provided by the manufacturer. Acceptable results between Col A and Col B are 0.005 or 5%, whichever is greater. When compared to the range provided by the manufacturer, both results (Col A and B) must be within the published range. If the values are greater than 0.005 or 5%, and/or either result from Col A or B is outside the published range, then the batch will be rejected and all cases re-tested.

REPORTING

1. All test results are recorded to three decimal places on worksheet(s).

2. Individual results from each case are averaged and round down to three decimal places.

3. Results are reported in grams of alcohol per 100 mL of blood.

4. Serum/Plasma results will be converted to whole blood with the conversion factor 1.22:1*

5. Concentrations of ethanol below 0.010 g/100ml will be reported as “Below minimum reporting       limit”

6. Worksheets and chromatograms for each sample are placed into the appropriate case folder along with a copy of batch sheet used in the calibration, and any other paperwork submitted with the case.  Each page will have the case number, and initials of the analyst.

7. Final reports, worksheets, chromatograms of both samples and standards are Technically and Adminstratively Reviewed.

8. The original worksheets, sequence list and chromatograms for calibrators and standards are placed in the Ethanol Controls folder labeled with the run date.  File this folder in an archives box.

9.  After the administrative review is completed, the analyst will complete the notarization of the report, and it will be sent to the customer.

* Measuring Blood Alcohol Concentration for Clinical and Forensic Purposes, AW Jones and Derrick Pounder, Handbook of Drug Abuse, S Karch, MD, 1998

## VII. 		  NOTES

Blood alcohol concentrations > 0.08 g/100 ml blood are prima facie evidence in operating under the influence violations.

Many volatile substances can be detected by this procedure.  The most common

volatiles in body fluids are ethanol, methanol, isopropanol, and acetone. All of these

substances can be separated from ethanol on the gas chromatograph.  It is the labs policy to indicate on reports to the customer the presence of other volatiles. The lab does not quantitate volatiles other than ethanol.

VIII. 	CASE DOCUMENTATION

A. CASE NOTES

The minimum information, which must be contained in the case notes are:

- Laboratory Identification Number

- Collection kit’s suspect/police information paperwork

- Run Data

- QC Data

- Comments/Results

All case notes, chromatograms and other data generated during analysis will bear the initials of the analysts and case number.  Addition notes may indicate the stopper color of submitted tubes, and if a collection kit was past the published expiration date at the time of collection.

B.  CASE FILE

The minimal information, which must be contained in the individual case file

consists of:

- The final report

- Any preliminary, supplementary or corrected reports

- Collection kit’s suspect/police information paperwork (if available)

- Worksheet(s)

- Evidence receipt

- Original chromatograms

- Technical and Administrative Review

- Uncertainty of Measurement

When estimating the uncertainty of measurement, all uncertainty components which are of importance shall be taken into account using appropriate testing procedures.  Bias of Calibrators is also acknowledged and examined as part of the UofM procedure.  Documentation, when applicable, will be retained by the Quality Manager.

- What is being measured:  Ethanol concentration in blood samples

- Traceability of is established by using NIST / Guide 34 traceable controls,  obtained by an approved vendor, and utilizing equipment calibrated to ISO 17025 standards by an accredited and approved vendor.

- The equipment used for determining ethanol concentration:

  - Headspace Gas Chromatograph- dual-column

    - Headspace- HP76954 S/N: IT80703190

    - Gas Chromatograph- HP7890 S/N: CN10809061

  - Hamilton Diluter – 600 Series: Microlab 600 - S/N:ML600FF8801

- The following components are recognized as potentially contributing to UofM:

  - Whole Blood control (Reproducibility)

  - Aqueous Controls (Bias of Calibrators)

  - Temperature (liquids and ambient)

  - Variation of time at room temperature

  - Uncertainty stated on COA’s

  - Matrix differences

  - Diluter:  samples, controls, calibrators, and internal standard

  - Stability of controls and calibrators

  - Staff variability

  - Headspace variability

  - Concentration

  - Internal Standard (n-propanol) stability and concentration

  - Stability

  - Instrument parameters

  - Instrument precision

  - Calibration model

  - Integration parameters / processing of data

- The following components are considered of significance:

  - Whole Blood Matrix Control - Type A: (reproducibility of assay)

  - Matrix - Type B: (5% administrative rule)

  - Calibrators- Type B from calibration certificate

  - Diluter for Internal Standard - Type B from calibration certificate

  - Diluter for Samples  - Type B from calibration certificate

  - Aqueous Controls - Type B from COA (to evaluate bias of calibrators)

- Data from controls and duplicates are tracked in a Microsoft Excel Spreadsheet.  Calibration certificate(s) of the Hamilton diluter, and COA’s of respective calibrators, QC standards, and whole blood controls are retained by the Quality Manager.  From these spreadsheets, and in particular the WBC’s, it can be determined that the Data is/is not of a normal distribution, skewed, and the mean and standard deviation calculated.  Additional graphs can also be created as warranted.  All values of uncertainty from individual components deemed significant (See E above) are converted to % uncertainty (See ASCLD/LAB Annex D AL-PD-3065 Ver 1.0).

- The following calculations are performed: the uncertainties of the six sources identified in section E, are individually squared, and totaled.  The square root is determined of the resulting sum, and this value is equal to the combined uncertainty or k.  The expanded uncertainty (K2) is calculated by doubling the uncertainty (Kx2).  (The method is identical to ASCLD/LAB Annex D AL-PD-3065 Ver 1.0).

___________________________________

√ (u1)2 + (u2)2 + (u3)2 + (u4)2 + (u5)2 + (u6)2    =K

K(2) = (Kx2) = reported UofM at 95.45% confidence interval

For reporting purposes, the lab will round up the K2 value for ease of use and understanding of case reports by our customers.

- The schedule to review the measurement uncertainty will be conducted annually or upon the addition or replacement of laboratory equipment, staff or other factors considered of significance.  The Quality Manager will retain calculations, verifications of spreadsheets, graphs, etc.

Reference: ASCLD/LAB Guidance on the Estimation of Measurement Uncertainty – Annex D. (ASCLD/LAB document: AL-PD-3065 Ver 1.0 22 May 2013)

# Text boxes

FAIL

.106

.100

Pass

Pass

.105

.101

Pass

FAIL

0.104 Result 2

0.096 Result 1

Pass

Pass

0.100 Standard